_____________________________________________________

LabWindows Data Acquisition Library README File

National Instruments LabWindows Data Acquisition

Library Software Version 4.0.1

Copyright 1991 National Instruments Corporation.

All rights reserved.

_____________________________________________________

This file contains information that is too recent to be included in

the printed LabWindows documentation.  The information in this file

pertains to the LabWindows Data Acqusition Library Reference Manual

(Part Number 320316-01).

1. EISA-A2000 board and bus mouse combination

The LabWindows interactive program may experience a system failure

when all of the following conditions are true:

- You have an EISA-A2000 running a SCAN or pre-trigger FRAME

oriented acquisition at 1M samples/sec with a buffer or frame

size of 25,000 scans or less.

- You have a bus mouse installed.

- You make a call to a User Interface Library function while

the EISA-A2000 is running.

To remedy this situation, use a serial mouse or buffer sizes greater

than 25,000 scans.

2. CTR_Simul_Op()

A new operation has been added to this function to allow multiple

counters to latch their counts simultaneously.  The description of

the mode parameter on page 8-22 of the LabWindows Data Acquisition

Library Reference Manual should read:

mode is the operating mode to be performed by this call.

mode = 1: Reserve counters specified in ctrList for simultaneous

start, restart, stop or count latch operation.

mode = 2: Perform a simultaneous start\restart on the counters

specified in ctrList.

mode = 3: Perform a simultaneous stop on the counters specified

in ctrList.

mode = 4: Perform a simultaneous count latch on the counters

specified in ctrList.  The counts may be retrieved one

at a time by subsequent calls to CTR_EvRead.

mode = 0: Cancel reservation of counters specifed in ctrList.

Note:  It is not necessary to call CTR_Simul_Op with mode set to 1

before calling CTR_Simul_Op with mode set to 4.	That is, it

is permissable to start two or more counters at different

times and still latch their counts at the same time.

3. New Error Code

The following error code should be added to page B-3 of the LabWindows

Data Acquisition Library Reference Manual:

-157   cannotLockBufErr      Unable to lock buffer inside the

interactive environment

4. MIO_16_F5_Calibrate()

The description of the calRefVolts parameter on page 3-3 of the

LabWindows Data Acquisition Library Reference Manual gives an incorrect

range of +6 V to +10 V, or -6 V to -10 V.  The correct range is +6 V

to +10 V.

Note:  You should always calibrate the ADC and the DACs after calibrating

the internal reference voltage.

5. New timebase available for the AT-MIO-16F-5

The DAQ_Start, SCAN_Start, and WF_Start calls now support a 5 MHz

timebase (200 nsec resolution) for the AT-MIO-16F-5.  To select this

timebase, pass a value of -1 for the timebase parameter of the DAQ_Start

or WF_Start calls, or the sampTimebase or scanTimebase parameters of

the SCAN_Start call.  The DAQ_Rate and WF_Rate will return a timebase

of -1 for intervals of 9 microseconds or less.

6. Notes on DMA Waveform Generation with the AT-MIO-16F-5

- The first point of the waveform buffer is not generated until the

second update pulse occurs.  That is, two update intervals will pass

before the waveform begins to appear at the I/O connector after the

waveform is started by a WF_Op, WF_from_Disk, or WF_Start call.

- An incorrect dacUpdateErr may be returned by WF_Check after a waveform

has finished if the waveform was loaded with an iterations parameter

that is greater than zero.  This will occur when the update interval is

relatively small compared to the speed of your computer.  For example,

on a 386 20 MHz PC AT computer with a computing index (CI) of 23.0,

WF_Check returns an incorrect dacUpdateErr after a waveform with

an update interval of 10 microseconds or less has completed its

its iterations.  In general, if a continuous waveform is error-free

at a given update interval on your computer, a terminating waveform will

also be error-free.  Thus, you can ignore a daqUpdateErr returned by

WF_Check when the status parameter is returned with a value of 1.

- When using double-buffered waveform generation with oldDataStop mode

enabled, bit 15 of the data points in the waveform buffer may be

altered by the driver.